## T.Minutes (2)
|  | MTI - HAYATABAD MEDICAL COMPLEX | [11 empty] | Doc. No. |  | HMC-PPD-F-05
|  | RECORD FORMAT | [11 empty] | Version No. |  | 0
|  | Minutes of Technical Evaluation & Advisory Committee Meeting | [11 empty] | Date |  | 2023-05-01
1.1 | Mobile X-Ray Machine:
Mobile X-Ray Machine was already selected in Cardiac Surgery Equipment.
Marking Sheet: 1.2: CT Scan (128 slices whole body computerized tomography scanner)
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 |  | 15
|  |  | 2: Product Evaluation; 6: Firm Evaluation; 14: Technical Evaluation |  | Financial Evaluation
| S. # | FIRM NAME | Compliance to HMC's Specification | Special Features | Product Certification | Product’s Local Performance Certificate | Intergrity Pact | Technical Staff |  | Work Station, Networking & Training |  | Previous Performance with HMC | Testing & Calibration Equipment | Post_ Warranty Series & Maintenance | Total Technical Marks |  | Total Financial Marks
[8 empty] | Engineers | Technical staff | Peshawar | Other Cities | Previous Performance with HMC
[8 empty] | (01 Prs) | ( 03 Prs)
[8 empty] | 4 | 02 each
|  | MARKS | 50 | 8 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 100 |  | 100
| 1 | M/S Medequips | 2: HMC Specifications requires GENERATOR & DETECTOR: High frequency type maximum power of at least 80kw or more whereas the firm quoted model have 72KW which is a major deviation.
| 2 | M/S Shirazi Trading | 2: HMC Specifications requires GENERATOR & DETECTOR: High frequency type maximum power of at least 80kw or more whereas the firm quoted model have 72KW which is a major deviation.
| 3 | M/S BIOS | 50 | 0 | 4 | 0 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 84
| 4 | M/S Hoora Pharma | 2: HMC Specifications requires GENERATOR & DETECTOR: High frequency type maximum power of at least 80kw or more whereas the firm quoted model have 72KW which is a major deviation.
| 5 | M/S Siemens Healthcare | 2: HMC Specifications requires GENERATOR & DETECTOR: High frequency type maximum power of at least 80kw or more whereas the firm quoted model have 75KW which is a major deviation.
| 6 | M/S Friends Traders. | 50 | 0 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 92
Passing Marks 70%
Marking Sheet: 1.3: DR Machine DIGITAL / FlOOR MOUNTED 550mA XRAY MACHINE
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 |  | 15
|  |  | 2: Product Evaluation; 6: Firm Evaluation; 14: Technical Evaluation |  | Financial Evaluation
| S. # | FIRM NAME | Compliance to HMC's Specification | Special Features | Product Certification | Product’s Local Performance Certificate | Intergrity Pact | Technical Staff |  | Work Station, Networking & Training |  | Previous Performance with HMC | Testing & Calibration Equipment | Post_ Warranty Series & Maintenance | Total Technical Marks |  | Total Financial Marks
[8 empty] | Engineers | Technical staff | Peshawar | Other Cities | Previous Performance with HMC
[8 empty] | (01 Prs) | ( 03 Prs)
[8 empty] | 4 | 02 each
|  | MARKS | 50 | 8 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 100 |  | 100
| 1 | M/S Medequips | 2: Non-responsive as the quoted model is not DR but a retrofit solution.
| 2 | M/S MEDCO | 2: The Mandatory Integrity Pact was not attached/provided in the quotation submitted by the firm, hence the firm's bid was non-responsive.
| 3 | M/S Fuji Film | 50 | 0 | 6 | 5 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 2 | 89
| 4 | M/S BIOS | 50 | 0 | 2 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 88
| 5 | M/S Hoora Pharma | 2: Non-responsive.
| 6 | M/S Siemens Healthcare | 50 | 0 | 4 | 1 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 1 | 82
| 7 | M/S Friends Traders. | 50 | 0 | 4 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 90
Passing Marks 70%
Marking Sheet: 1.4: Doppler Ultrasound Machine
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 |  | 15
|  |  | 2: Product Evaluation; 6: Firm Evaluation; 14: Technical Evaluation |  | Financial Evaluation
| S. # | FIRM NAME | Compliance to HMC's Specification | Special Features | Product Certification | Product’s Local Performance Certificate | Intergrity Pact | Technical Staff |  | Work Station, Networking & Training |  | Previous Performance with HMC | Testing & Calibration Equipment | Post_ Warranty Series & Maintenance | Total Technical Marks |  | Total Financial Marks
[8 empty] | Engineers | Technical staff | Peshawar | Other Cities | Previous Performance with HMC
[8 empty] | (01 Prs) | ( 03 Prs)
[8 empty] | 4 | 02 each
|  | MARKS | 50 | 8 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 100 |  | 100
| 1 | M/S Hoora Pharma | 46 | 0 | 4 | 0 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 2 | 78
| 2 | M/S Vertex Medical | 50 | 0 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 92
| 3 | M/S Medequips | 50 | 0 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 2 | 90
| 4 | M/S Friends Traders. | 50 | 0 | 6 | 4 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 90
Passing Marks 70%
Marking Sheet: 1.5: Digital Mobile X-Ray Machine with Battery Backup
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 |  | 15
|  |  | 2: Product Evaluation; 6: Firm Evaluation; 14: Technical Evaluation |  | Financial Evaluation
| S. # | FIRM NAME | Compliance to HMC's Specification | Special Features | Product Certification | Product’s Local Performance Certificate | Intergrity Pact | Technical Staff |  | Work Station, Networking & Training |  | Previous Performance with HMC | Testing & Calibration Equipment | Post_ Warranty Series & Maintenance | Total Technical Marks |  | Total Fnancial Marks
[8 empty] | Engineers | Technical staff | Peshawar | Other Cities | Previous Performance with HMC
[8 empty] | (01 Prs) | ( 03 Prs)
[8 empty] | 4 | 02 each
|  | MARKS | 50 | 8 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 100 |  | 100
| 1 | M/S Medequips | 50 | 0 | 4 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 2 | 88
| 2 | M/S Medifa | 46 | 0 | 2 | 4 | 2 | 6 | 4 | 5 | 5 | 2 | 2 | 4 | 82
| 3 | M/S MEDCO | 2: The Mandatory Integrity Pact was not attached/provided in the quotation submitted by the firm, hence the firm's bid was non-responsive.
| 4 | M/S BIOS | 50 | 0 | 2 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 88
| 5 | M/S Fuji Film | 50 | 0 | 6 | 1 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 2 | 85
Marking Sheet: 1.6: Single Plane Interventional Radiology Machine (Angiography)
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 |  | 15
|  |  | 2: Product Evaluation; 6: Firm Evaluation; 14: Technical Evaluation |  | Financial Evaluation
| S. # | FIRM NAME | Compliance to HMC's Specification | Special Features | Product Certification | Product’s Local Performance Certificate | Intergrity Pact | Technical Staff |  | Work Station, Networking & Training |  | Previous Performance with HMC | Testing & Calibration Equipment | Post_ Warranty Series & Maintenance | Total Technical Marks |  | Total Fnancial Marks
[8 empty] | Engineers | Technical staff | Peshawar | Other Cities | Previous Performance with HMC
[8 empty] | (01 Prs) | ( 03 Prs)
[8 empty] | 4 | 02 each
|  | MARKS | 50 | 8 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 100 |  | 100
| 1 | M/S Siemens Healthcare | 50 | 0 | 4 | 4 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 1 | 85
| 2 | M/S Friends Traders. | 50 | 6 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 3 | 97
Marking Sheet: 1.7: Color Doppler Trus (Trans-rectal Probe)
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 |  | 15
|  |  | 2: Product Evaluation; 6: Firm Evaluation; 14: Technical Evaluation |  | Financial Evaluation
| S. # | FIRM NAME | Compliance to HMC's Specification | Special Features | Product Certification | Product’s Local Performance Certificate | Intergrity Pact | Technical Staff |  | Work Station, Networking & Training |  | Previous Performance with HMC | Testing & Calibration Equipment | Post_ Warranty Series & Maintenance | Total Technical Marks |  | Total Fnancial Marks
[8 empty] | Engineers | Technical staff | Peshawar | Other Cities | Previous Performance with HMC
[8 empty] | (01 Prs) | ( 03 Prs)
[8 empty] | 4 | 02 each
|  | MARKS | 50 | 8 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 100 |  | 100
| 1 | M/S Vertex Medical | 50 | 0 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 92
| 2 | M/S Medequips | 50 | 0 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 2 | 90
| 3 | M/S BIOS | 50 | 0 | 4 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 90
| 4 | M/S Eastern Medical | 48 | 0 | 4 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 1 | 85
| 5 | M/S Friends Traders. | 50 | 0 | 6 | 6 | 2 | 4 | 6 | 5 | 5 | 2 | 2 | 4 | 92